## Slide 2
Fault locations
Identify impedance variation of connectors
Remove unwanted discontinuities
Simplify filter tuning

## Slide 3
Time domain reflectometry
0
I
I
L
0
I
Localization of the mismatch
Nature and amplitude of the mismatch
Echo measurement (similar to a radar)
Time domain measurement
Different times correspond to
different positions

## Slide 4: Matched Termination (ZL=Z0)
Z0
Ei
Open Circuit Termination (ZL=∞)
Ei
Ei
t
t
0
0
2T
Ei
T
ex (t)
ex (t)
Γ=0
Γ=1
Z0
Examples of DUTs

## Slide 5: Short Circuit Termination (ZL = 0)
0
t
2T
Ei
-Ei
Line Terminated in ZL = 2Z0
2Z0
t
2T
0
Ei
ex (t)
ex (t)
Γ=-1
Γ=
Examples of DUTs

## Slide 6: Line Terminated in ZL = Z0
Z0
0
t
2T
Ei
-
ex (t)
Γ= -
Examples of DUTs

## Slide 7
Examples of open and short circuit terminations (Screen captures from the 86100B)
Measurements

## Slide 8: Line Terminated with a Capacitor (ZL = C)
C
0
t
2T
Ei
τ = Z0 C
2Ei
t
2T
0
Ei
2Ei
τ =
L
Line Terminated with a Inductor (ZL = L)
ex (t)
ex (t)
Examples of DUTs

## Slide 11: Synthetic pulse with VNA (Time Domain Option)
Convenience: VNA’s outstanding dynamic range compared to the traditional TDR.
Time Response Transfer Function
The VNA operates the FFT of the Transfer Function, to obtain the Time Response.
FOURIER

## Slide 12
Z0
T
Γ
Z0
Step Response
Γ
t
0
0
Γ
Z0
Γ
Z0
Γ
t
0
0
Γ
t
0
0
2T
2T
1
-1

## Slide 13
2Z0
Γ
Z0
Step Response
Γ
t
0
0
Γ
Z0
Γ
Z0
Γ
t
0
0
Γ
t
0
0
2T
2T
1
-1
2T
-1
1
C
L

## Slide 14
0
I
I
2T
0
1
time
-1
Open termination

## Slide 15
0
I
I
2T
0
1
time
-1
Short termination

## Slide 16
0
I
I
2T
0
1
time
-1
-0.3
Z0/2 termination

## Slide 17
Capacitor

## Slide 18
Inductor

## Slide 20
The Time Domain Reflectometry in the VNA is a mathematical operation
(Chirp Z Transform)
1)
response to response to =
Step Response Pulse Response

## Slide 21
2) FFT algorithm requires symmetric frequency samples (with respect to the zero).
fmin ≠ 0 zero frequency sample MUST be extrapolated.
f
0
Δf
Δf
Δf
fmax
fmin = Δf
fmin depends by fmax and by the number of samples (number of points).
SET LOW FREQUENCY PASS before calibration!

## Slide 22
3) The slope of the simulated signal depends by fmax :
fmax1 > fmax2 > fmax3
RISE TIME
The spatial resolution is related to fmax .
The fmax is finite.
The Pulse Response isn’t but sinc(t) :

## Slide 23
4) Discrete sampling of frequency:
Tmax=1/2Df

## Slide 24
5) TIME DOMAIN FREQUENCY DOMAIN
(SYNTHETIC) (MEASUREMENT)
The FFT is influenced by the limited number of samples.
FFT
t/f
f/t
f0/ t0
t/f
f/t
BW
Δ
Rectangular window:
Δ= -13dB
BW depends by the number of samples

## Slide 25
5) TIME DOMAIN FREQUENCY DOMAIN
(SYNTHETIC) (MEASUREMENT)
The FFT is influenced by the limited number of samples.
FFT
t/f
f/t
f0/ t0
t/f
f/t
BW
Δ
Rectangular window:
Δ= -13dB
BW depends by the number of samples

## Slide 34
Can we measure the window used?
The pulse response to open or short loads is the FFT of the window being used

## Slide 35
Summary

## Slide 36
6) it’s possible to identify, in time domain, the part of response I’m iterested in.
GATING (time domain filter)
Example
Z0
Z1
Z0
Γ
DUT
Γ
t
Time Domain
(Pulse Response)

## Slide 37
Γ
t
Gating
Frequency Domain
Only DUT
Inconvenience:
Gating decreases signal’s power Lower SNR
General example: an attenuator will attenuate more than expected

## Slide 41
LOW PASS STEP
LOW PASS PULSE
BAND PASS
Used for band pass devices
Sign and amplitude of ρ isn’t always easy to understand
Used for gating

## Slide 43
Ten Steps for Performing TDR (Low Pass Step)
Set up desired frequency range (need wide span for good spatial resolution)
2. Under SYSTEM, transform menu, press "set freq low pass“
3. Perform one- or two-port calibration
4. Select S11 measurement *
5. Turn on transform (low pass step) *
6. Set format to real *
7. Adjust transform window to trade off rise time with ringing and overshoot *
8. Adjust start and stop times if desired
9. For gating:
  set start and stop frequencies for gate
  turn gating on *
  adjust gate shape to trade off resolution with ripple *
10. To display gated response in frequency domain
  turn transform off (leave gating on) *
  change format to log-magnitude *
* If using two channels (even if coupled), these parameters must be set independently for second channel
Notes: Slide 89
Here is a summary of how to perform TDR measurements. Without such a checklist, it is easy to overlook some of the more subtle steps, resulting in confusing or misleading measurements. A one-port calibration is all that is needed when characterizing connectors and the open, short and load standards. A two-port calibration is needed to characterize the reflection or line impedance of the thru standard.

## Slide 44: EXERCISES

## Slide 45
Fault location
Capacity variation (Folded cable)
C
Z0
Z0
T
Γ
Γ
Γ
t
t
2T
2T
LOW PASS
STEP RESPONSE
LOW PASS
PULSE RESPONSE
Inductance variation (Frayed cable)
L
Z0
Z0
T
Γ
Γ
Γ
t
t
2T
2T
LOW PASS
STEP RESPONSE
LOW PASS
PULSE RESPONSE

## Slide 46
Inductance variation (Frayed cable)
L
Z0
Z0
T
Γ
Γ
Γ
t
t
2T
2T
LOW PASS
STEP RESPONSE
LOW PASS
PULSE RESPONSE

## Slide 48: Impedance measurement as a function of Z
Z0
Z0
Z0
Z1
Z2
T1
T2
T3
T4
t
Γ
2T1
2T2
2T3
2T4
Z1 < Z0
Z2 > Z0

## Slide 49: Example: Microstripes
T1
T2
2T1
2T2
Nonideality:
The function doesn’t end at zero (attenuation, energy loss, reflections)
Second rise time is slower (dispersion, sharp corner are capacitive and cut high frequencies)
To have Z(x) you need to calculate the velocity

## Slide 50: Gating
DUT

## Slide 51: Microstipes
W
h
Impedances Various
Taper
Cross-Talk
50Ω
50Ω
25Ω
50Ω
100Ω
50Ω
100Ω
Differences?

## Slide 52
20 pF capacitor

## Slide 53
19 nH inductance

## Slide 55: APPLICATIONS

## Slide 66
Wire (0.5mm)
LHC TCTP collimator wire Measurement

## Slide 67
N. Biancacci et al., TCTP collimator meas., 2014
Coaxial transmission line properties
Matching in impedance measurement is critical.